c

# MALDIVES FOOD AND DRUG AUTHORITY

Ministry of Health

###### Male’, Maldives

Terms of Reference

CALIBRATION OF EQUIPMENTS USED AT NATIONAL HEALTH LABORATORY FROM A CALIBRATION CENTER ACCREDITTED TO ISO/IEC 17025

20th June 2018

- Introduction

National Health Laboratory (NHL) is a division of Maldives Food and Drug Authority (MFDA) with the main mandate of quality control, safety and hygiene verification of food and pharmaceuticals. NHL is divided into two main sections namely, Chemistry and Microbiology laboratories. The matrices tested   include water, food and swabs and brought from food establishments, factories and resorts and pharmaceuticals.

NHL has been accredited to the international standard ISO/IEC 17025 since December 2008. For the maintenance of its accreditation status it is compulsory to calibrate equipments from a calibration centre accredited to ISO/IEC17025.

- Scope and Objectives

The objective of this TOR is to find a competent calibration center / institute accredited to the international standard ISO/IEC 17025 to calibrate the equipments at NHL.

The List 1 attached includes the details of equipments that require calibration.

- Expected Outcomes

The chosen party is expected to:

- Calibrate the equipments in table 1 by sending technicians (for calibration of equipments that can be calibrated at NHL site) and by taking items to site of calibration (for items that cannot be calibrated at NHL.

- Issue calibration certificates with sufficient data for analysis and uncertainty calculations

- Place a sticker/label on the equipment with required details

- Requirements / Qualifications

  - The calibration centers intending to offer calibration services to NHL shall have the following qualifications and requirements.

- Must be a calibration center accredited to the standard ISO/IEC 17025.

- Calibration certificates/reports must be issued, in paper, after calibration, bearing the accreditation logo.

- Calibration certificates  must at least contain the following information:

  - Description of the equipment

  - Serial No.

  - Identification No/Manufacturer and Model no.

  - Calibration Date

  - Location of calibration

  - Test conditions:

  - Reference standard used

  - Method of calibration

  - Uncertainty associated with the activity

- A label/sticker should be placed on the equipment after calibration with the details:

  - LOGO of the calibration center

  - A reference number

  - Identification number of the equipment

  - Calibrated by

  - Date of calibration

- Calibration certificates should be received to NHL in at least a month’s time.

- The calibration service provider must have at least 3 years’ of experience and must be currently engaged in providing the service.

  - Documents to prove conformance to the points numbered 4.1 a( of above criteria should be submitted along with proposals and proof to show conformance to the rest of the criteria should be submitted either by submitting relevant documents or statements in the proposals. Submissions without these documents and statements will be cancelled. Proposals submitted by parties not conforming to the above criteria will be cancelled.

- Quoting and required documents

  - Prices must be quoted for the calibration service of all items in the table 1. Proposals and quotations which do not quote for all items will be cancelled.

  - The quotations must indicate clearly the following charges in in Maldivian Rufiya (MVR) or US Dollars ($):

  - Rate of calibration/consultancy charges for each type of equipment and the total cost

  - Postage costs involved in sending any required equipments to the calibration center

  - Postage costs in returning the equipments to NHL, Maldives

  - Return airfare for technicians for the calibration to arrive to Maldives and their return

  - Accommodation charges for the calibration technicians for the period of stay in Maldives

  - No. of days required for the calibration done in NHL and posted items

  - Charges for any freight of devices/ equipments / solutions/ standards required for the calibrations that would be brought in

  - Quotations /proposals not stating the above will be cancelled. Any other charges or incidental charges that would incur must be included. Any charges not mentioned in the quotation will NOT be borne by NHL during the payment process.

  - Equipments that cannot be calibrated in the laboratory site should be identified in the quotation.

  - Arrangements that need to be made, standards and apparatus that need to be available from NHL such as required environment conditions must be stated.

  - In addition to the quotation interested parties should submit:

        - A copy of the accreditation certificate to ISO/IEC 17025

        - A sample report of calibration

        - Brief profile

        - List of clients/ laboratories that had sought calibration services in the past 3 years.

        - Payment conditions

        - Any other terms and conditions subject to the process

        - Contact details

  - Proposals/ quotations not accompanying the above documents/ information except for 4.2.(f) will be cancelled.

- Duration: All calibrations are to be completed before the 30th November 2018.

- Selection:  An appropriate party shall be chosen by evaluating the quotations/proposals using the following grading system. The party which gets the maximum number of points shall be awarded.

- Price: 80 (Eighty) The highest point will be given to the least quoted party and the rest will be graded by ratio.

- Experience: 20 (Twenty) points; External documents stating the duration of providing calibration services. The longest duration gets the highest point and the rest will be graded by ratio.

- Contact:  Thooma Adam/ Maldives Food and Drug Authority / Phone: 00 960 3014346

- Deadline for submission:

Interested parties are to submit signed quotations and other required documents on or before the official time of 28th June 2018 to

Maldives Food and Drug Authority, Ministry of Health,

Roashanee Building,

Sosun Magu,

Male’ 20184

Republic of Maldives.

Electronic versions of the quotations and documents can be emailed to nhl@health.gov.mv before the above due date. Late submissions cannot be considered under any circumstances.

TABLE 1 -NHL CALIBRATION LIST 2018

| # | Item | Manufacturer/Model No. | Serial/ Number | Lab Identification Number | Performance Range/Capacity | Resolution | Location/Lab | Accuracy |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Digital burette (50mL) (New) | Brand | 8P6165 | NHL/CH-WC/EQ219 | 50 ml |  | Water chemistry |  |
|  | Thermometer | Alla France | - | NHL/CH-WC/EQ172 | -10 - +110 0C |  | Water chemistry |  |
|  | MK II Chloride analyzer 926 | Sherwood | 10639 | NHL/CH-WC/EQ057 | - |  | Water chemistry |  |
|  | BOD Incubator | Hach | 06092708 | NHL/CH-WC/EQ079 | 2-4°C |  | Calibration Room | 200c |
|  | Refrigerator | Whirlpool | - | NHL/CH-WC/EQ077 | Chiller 2-8°C Freezer <18 | +/- 2 | Water chemistry | 1 0C |
|  | Fume hood | ESCO | 2009-39431 | NHL/CH-GN/EQ146 | - |  | Water chemistry |  |
|  | Turbidimeter | Hach 2100N | - | NHL/CH-WC/EQ058 |  |  | Water chemistry | 1NTU |
|  | COD Reactor | Hach USA | 980400017502 | NHL/CH-WC/EQ064 | - | - | Water chemistry |  |
|  | pH meter | EUTECH | 1569174 | NHL/CH-WC/EQ109 | - | - | Water chemistry | 0.01 PH units |
|  | pH meter (New) | Hach (New) HQ11D | 150900016173 |  |  |  | Water chemistry | - |
|  | Dispenser (New) | BRAND GMBH + CO KG | 07G 94025 | NHL/CH-WC/EQ205 |  |  | Water chemistry | 0.12% |
|  | Laboratory Oven (NEW) | Thermo scientific Hera therm OGH 100 51028143 | 42029579 | NHL/CH-GN/EQ223 | 50˚C-250˚C | 2°c | Bottle wash& distillation |  |
|  | Micro Pipette (NEW) | Fisherbrand | CJ12069 | NHL/CH-FC/EQ217 | 20ul to 200ul | 1ul | Food chemistry |  |
|  | Micro Pipette (multi-channel) (NEW) | Transferpipette-12 | 11Y3003 | NHL/CH-WC/EQ043 | 20ul to 200ul | 1ul | Food chemistry |  |
|  | Micro Pipette (FINNPIPETTE) (NEW) | Thermo scientific | MW 16085CE | NHL/CH-FC/EQ224 | 0.5-5ml | 0.5mL | Food chemistry |  |
|  | OMT Oven | Gallenkamp OMT 075.XX 2.C | SG98/02/314 | NHL/MB-FM/EQ006 | 55°C±2 | 0.1°C | Bottle Wash | 55°C±2 |
|  | Stomacher | Seward 400 | 47290 | NHL/MB-FM/EQ075 | 30,60,120 sec ±2 | ± 1 | Food Microbiology | 30,60,120 sec ±2 |
|  | Incubator | Sanyo MIR 262 | 40608373 | NHL/MB-WM/EQ004 | 44°C±0.5 | 0.1°C | Water Microbiology | 44°C±0. |
|  | Biological Safety Cabinet II (New) | Thermo Scientific | 300116932 | NHL/MB-MB/EQ113 | - | - | Media Room | - |
|  | Biological Safety Cabinet (New) | BioBase BSC-200011A2-X | BSC20A 1609008 | NHL/MB-FM/EQ114 | - | - | Food Microbiology | - |
|  | Dispenser (New) | Assimat | - | NHL/MB-MR/EQ110 | 1 to 10 ml | - | Media Room | 9ml |
|  | Pipette (New) | Finn pipette F2 | MW16984 | NHL/MB-MR/EQ115 | 1 to 10 ml | - | Media Room | 10ml |
|  | Incubator | Gallenkamp IPR 075XX 1.5 | SG97/04/367 | NHL/MB-FM/EQ002 | 44°C±0.5 | 0.1°C | Water Microbiology | 44°C±0.5 |
|  | Ultra-Low Temperature Freezer | Sanyo MDF-U73V | 61016921 | NHL/MB-VL/EQ039 | < -18°C±2 | 1°C | Molecular Biology | < -18°C±2 |
|  | Analytical Balance | JD 200-3 | A191658 | NHL/CH-PC/EQ198 | Up to 200g | 0.001g | MTG Laboratory | 0.001g |
|  | Analytical Balance | Sartorius | CPA225D | NHL/CH-GN/EQ149 | 0.01mg – 220g | 0.01mg | Sensory Evaluation 1 |  |
|  | Immersion thermometer (Reference) | Brannan | - | NHL/CH-FC/EQ042 | 0-210°C | 1°C | Sensory Evaluation 1 |  |
|  | Weight Set (12 pieces) | Troemner | - | NHL/GN-CH/EQ159 | From 100g to 5mg | - | Sample Receiving | 20mg, 50mg, 100mg, 500mg, 2g, 5g, 10g, 50g, 20g, 100g |
|  | Weighing Balance | Sartorius | 38110175 | NHL/CH-GN/EQ 151 | 0 to 420 g | - | Sample Receiving |  |
|  | Refrigerator | Whirlpool | - | NHL/CH-GN/EQ145 | - | - | Chemical Store | 2-8 0C |